Sears

SIMPSON	260	METER

REPAIR	PROGRAM

DIVISION	92

SOURCE	260

MAY,	1976

This booklet is designed as a simple tr shooting guide to repair a Simpson 260 meter in your Service Center. Al repairs are NOT covered, only the easy repairs that do not require major recalibration. U after following the enclosed procedure the problem has not been resolved, return the meter to the authorized repair source for your Territory.

A complete schematic an parts location list is included for your convenience for Model Series	6P.

Repair parts are available from your Territorial Repair Parts Distribution Center. These par are listed on ultrafiche in Division 92,  Source 260

SERVICE PROCEDURE - SIMPSON  260-3-4-5-SP-6-6P

THE  FOLLOWING	TESTS	CAN	BE

)	PERF ORM ED	TO		DE T ERMINE DEFECT S	AND	D EFECTIVE

COMPONENTS:

Set the instrument on the Rxl range, short the test leads together and adjust the zero ohms for zero indication on the meter. If the pointer falls short of zero, REPLACE THE 1.5 VOLT "D" SIZE BATTERY. This should restore normal operation.

Next, set the 260 to the RxlOK range and adjust for zero ohms. If the pointer will not reach zero, change the

_	9 volt transistor radio type (NEDA 160

3. If the pointer does not deflect. in Steps 1 and 2, remove the test leads from their jacks and short the "common" and "plus" jacks together with a paper clip. The 260 should still be set on a resistance range while making this check. If the pointer moves, REPLACE THE DEFECTIVE  TEST   LEADS.

4. If the po	did not deflect in Step 3

_	re move	the check the i

.	fuse is accessible through the battery compartment.

5. If Steps 1 to 4 do not loc  the trouble, or if the problem involves a single range only (such as Rxl or 1 MA) then an odor or visual test should be made to loc a burned resistor. The location of the components are as shown on the drawings for the ranges most likely to cause difficulty.

NOTE

1: If a resistor burns out and causes the printed circuit board to become crystalized,  the entire board must be

replaced because a degree of continuity is often set up across this area. This causes inaccuracies in the instrument and usually another breakdown in a short time.

NOTE

2: With the instrument out of the case be sure to check the switch decks for smoked areas or bum marks between contacts. If there is evidence of burn or arc marks, return unit to your authorized repair center.

NOTE

Switch replacement, C OMPLE TE calibration, meter movement replacement, or repair is considered a factory repair. Send the complete unit to your Authorized Repair Station.

OVERLOAD RELAY ADJUSTME

NOTE

Field adjustments of the Overload Relay Leaf Switch causes intermittent or erratic ope	of the unit. Replacement is recommended.

The mechanical adjustment of the relay is pe	as follows;

Refer to the drawing to locate the adjustment screw for the relay latching function. This can be adjusted with a small jeweler's screwdriver while the reset button is pushed in.

With the reset button depressed, tum the screw out until it does not catch on the relay. No more than one or two turn should be required.

Short the common and plus terminals and set the range and function switch on RxlOO and plus DC. This will provide an indication on the meter when the relay contacts are open or closed. Meter will read full scale when contacts are closed.

Now push the reset button and turn the screw clockwise until it just catches on the relay. This is to be done in small increments while allowing the button to spring in and out until it snaps into the hole.

The mechanism should be operated two or three times to be certain the latching an switch function is consistent. This can be accomplished by finger continuity between shunt R22 and the

This should cause the relay to release as indicated by the button "popping out".

NOTE

Finger contact at various points in the tester can cause the re to release. Do not mistake accidental finger contact with the circuitry with a defective relay or its adjustment.

positi	terminal of the 15 volt battery.	ANY REPAIR NOT COVEREDIN THESE

Push in the resetbutton and then connect	PR OCEDURES	IS	CONSIDERED    A

the two points together wi	the fingers.	FACTORY   REPAIR.

TROUBLESHOO	CHART,	I    260-6P

SYMPTOM	PO	DEFE

Rx1 will not zero	a.	Defe	"D" with short	in-		Battery

CURE

Replace

NOTES

put.

Ope	resistor

Replace

Defe	range switch

Rx1Ok will not	a.	Defe	battery zero with shorted

input.

c.

Replace	a.

260-6,6P = R16,17,19,21.

Se   switch re-

placement

.

.

12=9v

Open resistor

Defective range switch SWl

Rx100 will not	a.	Open resistor zero with shorted

b.	Replace	b.

c.

Replace	a.

NEDA 1604

R18

See switch re- placement

R18

input

Defective switch SW-1

Se   switch re- placement

| 7.5 VAC range rea approx. 1/2 | a. | Ope diode Dl or D2 | a. | Replace | a. | Se calibration proc |
| --- | --- | --- | --- | --- | --- | --- |
| the actual voltage |  |  |  |  |  |  |
| Out-of-tolera | a. | Calibration | a. | Recall- | a. | Se calibration |
| indication at full sc on 2.5 VAC | b. | req Dl or D2 de- | b. | brat• Replace |  | proced |
| range |  | fective |  |  |  |  |
|  | c. | Defe call- |  | Replace | c. |  |
|  |  | bration control |  |  |  |  |
|  |  |  |  |  |  | R22,25 |
|  | d. | Defe re- | d. | Replace | d. |  |
|  |  | sistor |  |  |  |  |
|  |  |  |  |  |  | --- -,--- R23,24, |
|  |  |  |  |  |  | 26 |

| 6. | SYMPTOM AC voltage range(s) | a. | PO DEFE Defe re- | a. | CURE Replace | a. | NOTES |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | defective |  | 1i1ton |  |  |  |  |
|  |  |  |  |  |  |  | --- -,-- R27 thru R30 |
|  |  | b. | Defe range |  |  | b. | S. switch re- |
|  |  |  | switch SW-1 |  |  |  | placement. |
| 7. | DC current ranges de- | a. | Ope shunt re- sistor | a. | Replace | a. |  |
|  | fec |  |  |  |  |  | Rll |
|  |  |  |  |  |  |  | thru R15. |
|  |  | b. | Defective range |  |  | b. | S. switch re- |
|  |  |  | switch (SW-1) |  |  |  | placement. |
| 8. | DC voltage range(s) de- | a. | Ope multiplier re1i1ton | a. | Replace | a. |  |
|  | fec |  |  |  |  |  |  |
|  |  |  |  |  |  |  | R4 thru RI |
|  |  | b. | Defective range |  |  | b. | S. switch re- |
|  |  |  | switch (SW-1) |  |  |  | placement. |

TROUBLESHOO

CHART, MODELS 260-3 THRU 260-6P

CALIBRATION CHECK

DC VOLTAGE RANGES can be checked for accuracy  by  usi		a battery of prope voltage for the range being checked while monito	the	battery	voltage	being applied across the test leads with another 260 or equivalent meter.

RESISTANCE RANGES can be  checked for accuracty by measuring a known value of resistance wi   the instrument.

AC VOLTAGE CALIBRATION

When the diodes or resistors in the bridge recti	have been replaced, the 250 AC and

2.5 VAC ranges must be calibrated.

250 VAC RANGE CALIBRATION

Using the line voltage from an AC outlet as a calibrating voltage source and an accurate AC monitoring voltmeter, connect the test leads an  adjust	or

R22 1	for the same reading as the monitoring voltmeter.

VAC RANGE CALIBRATION

Using an oscillator set to 1 volt RMS output at 60 hertz as read on a monitoring AC voltmeter. Set the 260 range switch to the 2.5 VAC range and connect the test leads to the

monitored  1 VAC  output of the oscillator.

-

Adjust      ,---- __	__  ,	R25 :

for the same reading as the monitoring voltmeter. This completes the AC calibr ation. SEE NOTE ON FOLLOWING   PAGE.

Simpson 260-6 & 6P Parts location

NOTE:	260-6  DOES  NOT  HAVE PROTECTION  BOARD

RANGE SWlH--,--------------------,-------------------7	1

I

1	It	1/	I

1	�v.V.

I	I	I

2     �v.	2	2

3

3    a".J';.

2

3

121

4    lOY.	4	4

5     jl;.5V.	5

6     f.C12.,MA.	6

7      IOMA.	7

4	200l

12

5

I   .10..MA.

lMA.

IX 1

11       Ix 10	112	11

12

I

117

12      Ix 10.0

9Y

ZERO OHMS

f   I

I	1.

I    -D.C..  I

I	/	I

2   +D.C.     /

I

2

3	A.C.

+Il	IODOY.	+	COM_MON	+IOl.

- SWl IS IN 50CW10 VOLn POSITION.

� SWl IS .. +D.C.

K• 1,00 OHMS.

MI0.•\00,00 OHMS.

SCH EMATIC DIAGRAM MOD EL  260-6

I	ns 1u

I

OM·l	132	D7

-

I	:

:

I

I

c2

DI	.002

l

I

I

1	Y2	1	i

I

I

I.ANGE SWITCH--.----- --- -

SW·I

1  =	'	..

2!�Y.	2

3   SOY.	3

JIAMl'S.

___   - -- ---  - - -   -  -  ___ _J

S    UV.

ov.

SODll

-

IGOMA.

I  10MA..

....

IMA.

10   IX I

11  IX10

4

113

22.S

10

m

12

4

s

I

•

DI

117	Ill

II

mn

122

+IV

SW

FUNCTION SWl1C

I -D.C. • I	I

2 + D.C.

3   A.C.

..

7.S

....

110

\Gmv.

ac.

--- ·

I	I•

I

2

3

ch)

,I

OUTPUT

I

2

3

� +IOA.

Ill

-IOA.

SCHEMATIC DIAGRAM MODEL 260-6P